## Slide 1
MS Thesis Presentation
Software Defined Radio Based Vector Network Analyzer for Characterization of RF Sensors and Circuits
Presented by: Ali Raza
Supervisor: Dr. Rashad Ramzan
Department of Electrical Engineering
National University of Computer Emerging Sciences Islamabad

## Slide 2: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 3: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 4: Vector Network Analyzer (VNA)
Measure the scattering parameters of electrical networks
Needs calibration for de-embedding the measurements

## Slide 5: Software Defined Radio (SDR)
Replaces traditional hardware with programmable hardware
Can be used for various applications by software reconfiguration

## Slide 6: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 7: Problem Statement
VNA is a very expensive instrument
Most essential instrument at RF laboratories
Low cost SDR can replace commercially available high cost VNAs for trivial applications

## Slide 8: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 9: Objectives
Pluto SDR programing for desired reception
Processing the data in MATLAB/Simulink
Design and fabrication of a TRL calibration kit
Implementation of TRL algorithm
S-parameter measurement using Pluto SDR
Application of TRL calibration on the measured S-Parameters
Verification of measurement results with bench top reference VNA

## Slide 10: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 11: Background
Joseph Mitola proposed the concept of SDR in early 90s [4]
Over the years SDR achieved great flexibility
SDR commercially used since last decade
SDR has been widely extended to different fields of electronics
[4]: J. Mitola, "Software Radios: Survey, Critical Evaluation and Future Directions," IEEE Aerospace and Electronic Systems Magazine, vol. 8, no. 4, pp. 25 - 36, 1993

## Slide 12: Some of Commercially Available SDR
ADALM-PLUTO SDR
BladeRF x 115
HackRF
USRP B200

## Slide 13: Some of Commercially Available SDR
 | ADALM-PLUTO | HackRF | BladeRF x 115 | USRP B200
Operating Frequency | 325 MHz – 3.8 GHz (70 MHz – 6 GHz with software modification) | 1 MHz - 6 GHz | 300 MHz - 3.8 GHz | 70 MHz – 6 GHz
Bandwidth | 20 MHz | 20 MHz | 28 MHz | 56 MHz
Duplex | Half or full duplex | Half duplex | Hal or full duplex | Half or full duplex
BUS | USB 2.0 | USB 2.0 | USB 3.0 | USB 3.0
ADC/DAC Resolution | 12-bit | 8-bit | 12-bit | 12-bit
Million Samples per Second | 61.44 MSPS | 20 MSPS | 40 MSPS | 61.44 MSPS
Cost | $ 249 | $ 300 | $ 650 | $ 966

## Slide 14: ADALM-PLUTO SDR
Portable self-contained RF learning module
Cost effective experimentation platform
One transmitter and one receiver
Receiver gain: 0 – 72 (dB)
Transmitter gain: -87.5 – 0 (dB)
MATLAB and Simulink support
USB powered

## Slide 15: Scattering Parameters
Describes reflection and transmission responses of a DUT
N-ports DUT have N2 elements in scattering parameters matrix
S11 represents reflection from port 1
S22 represents reflection from port 2
S21 represents transmission through port 2
S12 represents transmission through port 1

## Slide 16: Calibration
Removing errors from measurement planes
Different types such as SOLT, LRM and TRL
Removes systematic errors which occur due to the non-ideal components in the test setup

## Slide 17: Classical VNA Architecture

## Slide 18: Coupling Signals
Directional coupler separates the incident and reflected signals at the input port of the DUT
They have a through path and a coupled path
An amount of power travelling along through path is coupled
Coupled power is defined by the coupling factor

## Slide 19: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 20: Proposed SDR Based VNA Architecture

## Slide 21: Interfacing with MATLAB
A C++ code will interface the Pluto SDR with MATLAB
Saves the digital data of incident signal as a reference
Saves the digital data of reflected and transmitted signal as response
Ratio of the saved data to calculate the S-parameters

## Slide 22: Pluto SDR Before and After RF Shielding

## Slide 23: Loopback Mode

## Slide 24: Measurement Setup

## Slide 25: Directional Coupler
The directional coupler (CPL-5230-10-SMA-79) used is a product of Midwest Microwaves
It possesses inherently low insertion loss and VSWR with high directivity
It is an Ultra-Wideband device covering frequency range from 0.5 GHz to 18 GHz
Provides a 10 dB coupling output, 0.8 dB insertion loss, 15 dB directivity and 1.5 VSWR

## Slide 26: MATLAB Code Flowchart

## Slide 27: MATLAB Code Flowchart

## Slide 28: System Parameters
SDR transmitter and receiver setting:
Frequency (500 MHz – 3800 MHz)
Frequency step (6.44 MHz) (512 points)
Tx/Rx baseband sample rate: 100 KHz
Rx samples per frame: 5000
Generated sinewave:
Power: 0 dB
Frequency: 5 KHz
Samples per frame: 5000

## Slide 29: TRL Calibration
Through Reflect Line (TRL) calibration
Recommended for 1 to 10 GHz
Accuracy depends upon the quality of the fabricated kit
As the name implies, there are three standards
Through: Ideally it has a length close to zero, zero loss and zero reflection
Reflect: Any length, as long as it has a high reflection
Line: Recommended electrical length is ¼ λ

## Slide 30: Fabricated TRL Kit
Substrate used: ROGERS RO4360G2
εr: 6.15
Thickness: 1.524 mm
Dissipation factor: 0.0038

## Slide 31: S-Parameters of “Through” Standard

## Slide 32: S-Parameters of “Reflect” Standard

## Slide 33: S-Parameters of “Line” Standard

## Slide 34: TRL Calibration Algorithm
Determines deviations from the ideal measurements
Deviations occur as a result of coaxial to micro strip transition
Two error boxes are introduced in the measurement reference plane
Obtain the parameters of these error boxes and remove their effect from the subsequent measurements.

## Slide 35: Parameters of Error Boxes
Cascade of error box A and B by connecting the measurements planes 1 and 2 together using “through” standard
Terminating the measurement planes by connecting them with the “reflect” standard
Cascade of the error box A and B and the “line” yields the third
Finding the scattering parameters of all these combination it is possible to solve the individual parameters of the two error boxes and remove their effect from the measurements

## Slide 36: TRL Calibration Algorithm
The calibration is performed on the basis of the TRL algorithm presented in the Journal Paper [1]
The algorithm performs the TRL calibration in two main steps
1st the S-parameters of the TRL standards and the DUT are measured
In the 2nd step the effects of the TRL text fixture are removed from the DUT measurements
[1]: G. F. Engen and C. A. Hoer, "Thru-Reflect-Line: An Improved Technique for Calibrating the Dual Six-Port Automatic Network Analyzer," IEEE Transactions on Microwave Theory and Techniques, vol. 27, no. 12, pp. 987 - 993, December 1979

## Slide 37: TRL Calibration Algorithm
Step 1
Four matrices having s11, s21, s12, and s22 are created for the “line” and “through” standards and the DUT
Two matrices having s11 and s22 are created for the reflect standard
To calculate the two error boxes A and B presented in the article, we first find the wave cascading matrices for line, thru and the DUT represented as Rd, Rt and Rm in the article respectively

## Slide 38: TRL Calibration Algorithm
Step 2
The calibrated s-parameter matrix of the DUT is found using the following equation
By finding all the unknown values in the equation above as explained in the TRL algorithm article we get the calibrated S-parameter matrix of the DUT.
All the unknown values are found by simplifying the mathematical equations explained in the article

## Slide 39: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 40: Results
These results were carried out with the following settings
PLUTO SDR Settings | 
Frequency | 0.5 GHz – 3.8 GHz (512 points)
Frequency step | 6.44 MHz
Power | 30 dBm
Samples per frame | 5000
Moving average points | 3
VNA (MS46122B) Settings | 
Frequency | 0.5 GHz – 3.8 GHz (512 points)
Frequency step | 6.44 MHz
Power | 23 dBm max.
IFBW | 100 KHz
Averaging | Off

## Slide 41: DUT for Measurements
Single Stub
Double Stub
Patch Antenna

## Slide 42
TRL Calibration on SDR Measurements

## Slide 43: Single Stub S11 Parameter
Chart: Frequency (GHz)

## Slide 44: Single Stub S21 Parameter
Chart: Frequency (GHz)

## Slide 45: Double Stub S11 Parameter
Chart: Frequency (GHz)

## Slide 46: Double Stub S21 Parameter
Chart: Frequency (GHz)

## Slide 47: Patch Antenna S11 Parameter
Chart: Frequency (GHz)

## Slide 48
SDR and VNA Comparison

## Slide 49: Single Stub S11 Parameter
Chart: Frequency (GHz)

## Slide 50: Single Stub S21 Parameter
Chart: Frequency (GHz)

## Slide 51: Double Stub S11 Parameter
Chart: Frequency (GHz)

## Slide 52: Double Stub S21 Parameter
Chart: Frequency (GHz)

## Slide 53: Patch Antenna S11 Parameter
Chart: Frequency (GHz)

## Slide 54: Pluto SDR Measurement Setup
Measurement of Single Stub Resonant Sensor
S21 Measurement Setup
S11 Measurement Setup

## Slide 55: Pluto SDR Measurement Setup
Measurement of Double Stub Resonant Sensor
S21 Measurement Setup
S21 Measurement Setup

## Slide 56: Pluto SDR Measurement Setup
Measurement of Patch Antenna
S11 Measurement Setup

## Slide 57: VNA Measurement Setup
Measurement of Single Stub Resonant Sensor

## Slide 58: VNA Measurement Setup
Measurement of Double Stub Resonant Sensor

## Slide 59: VNA Measurement Setup
Measurement of Patch Antenna

## Slide 60: Presentation Outline
Introduction
Problem Statement
Objectives
Background
Design Philosophy and Implementation
Measurement Results
Conclusion

## Slide 61: Conclusion
In this thesis we have presented a low cost SDR based VNA
The operating frequency of the proposed VNA is 0.5 GHz to 3.8 GHz
The SDR based VNA is calibrated using a TRL calibration kit
The performance of the proposed VNA has been compared with a commercial VNA from Anritsu
A quite good resemblance is observed between the results obtained from SDR based and commercial VNA
Thus, it can be used as a cheap instrument in the research laboratories and educational institutes for characterization of RF devices

## Slide 62: References
G. E. a. C. Hoer, "Thru-Reflect-Line: An Improved Technique for Calibrating the Dual Six-Port Automatic Network Analyzer," IEEE Transactions on Microwave Theory and Techniques, vol. 27, no. 12, pp. 987 - 993, 1979
A. Estrada, "The vector network analyzer - an essential tool in modern ATE measurements," IEEE Instrumentation & Measurement Magazine, vol. 15, no. 4, pp. 22-26, 2012
D. Rytting, "ARFTG 50 Year Network Analyzer History," in 2008 71st ARFTG Microwave Measurement Conference, Atlanta, 2008
J. Mitola, "Software Radios: Survey, Critical Evaluation and Future Directions," IEEE Aerospace and Electronic Systems Magazine, vol. 8, no. 4, pp. 25 - 36, 1993

## Slide 63
Thank you